### 1. Purpose of the Procedure

- Measure Corneal Curvature: To determine the curvature of the anterior corneal surface (specifically the central 3mm zone).

- Assess Astigmatism: To quantify the amount and axis of corneal astigmatism and differentiate it from internal (lenticular) astigmatism.

- Contact Lens Fitting: To determine the Base Curve required for a contact lens so it fits the shape of the eye properly (not too tight or too loose).

- Detect Abnormalities: To identify irregular corneal surfaces, such as in Keratoconus (where the mires appear distorted) or corneal scarring.

- Biometry: To calculate the power of the Intraocular Lens (IOL) needed before cataract surgery.

### 2. Equipment Required

- Keratometer (Ophthalmometer): The standard manual device (e.g., Bausch & Lomb type) is most common in diploma courses.

- Occluder: To cover the eye not being tested.

- Alcohol Swabs/Chin Rest Papers: For hygiene.

### 3. Detailed Procedure (Bausch & Lomb Type)

#### Part A: Setup and Focusing

- Focus the Eyepiece:

  - Why? To ensure the examiner's own vision doesn't affect the measurement (accommodation).

  - How: Hold a white piece of paper or an occluder in front of the objective lens. Turn the eyepiece counter-clockwise until the cross-hair (reticle) is blurry. Then, slowly turn it clockwise until the cross-hair is sharp and black. Stop immediately once it is clear.

- Patient Positioning:

  - Adjust the chin rest and stool height so the patient is comfortable.

  - Align the patient's outer canthus (corner of the eye) with the black marker on the side of the instrument.

  - Instruct the patient to place their chin in the cup and forehead firmly against the rest.

  - Occlude the left eye to test the right eye first.

#### Part B: Alignment

- Gross Alignment:

  - Turn the instrument on.

  - Look from the outside (side of the machine). Raise or lower the instrument until the light from the keratometer reflects roughly in the center of the patient's cornea.

- Fine Alignment:

  - Look through the eyepiece. You will see three circles (mires).

  - Use the joystick to center the cross-hair inside the bottom-right circle.

#### Part C: Measurement

- Focus the Mires:

  - Use the focusing knob (usually the large outer ring or joystick twist) to make the bottom-right circlesingle.

  - Tip: If the circle looks doubled, the instrument is out of focus. Adjust until it is sharp and single.

- Find the Axis (Principal Meridians):

  - Rotate the entire tube of the instrument (using the handles) until the horizontal lines of the plus (+) signs and the vertical lines of the minus (-) signs are continuous and aligned with their spurs.

- Measure the Horizontal Meridian (Primary):

  - Turn the left measuring drum to move the plus (+) signs.

  - Adjust until the two plus signs overlap perfectly to form a single cross.

- Measure the Vertical Meridian (Secondary):

  - Turn the right measuring drum to move the minus (-) signs.

  - Adjust until the two minus signs overlap perfectly to form a single line.

### 4. Interpretation of Mires

- Clear/Regular: If the circles are round and sharp, the corneal surface is regular.

- Distorted/Wavy: If the circles look wavy or broken, it indicates irregular astigmatism or dry eye. Ask the patient to blink; if it clears, it was dry eye. If it stays wavy, it may be corneal scarring or Keratoconus.

### 5. Documentation (Recording Method)

Record the Power (in Diopters) and the Axis (in degrees) for both the horizontal and vertical meridians.

- Format:

  - Method 1 (Power/Axis): Record the two readings separately.

    - OD: 42.50 @ 180 / 43.50 @ 090

  - Method 2 (Flat/Steep): Identify the flatter curve (lower number) and steeper curve (higher number).

    - OD: 42.50 / 43.50 @ 090

- Mires: Note the quality.

  - Mires: Clear.

- Astigmatism: Calculate the difference (cylinder).

  - Example: 43.50 - 42.50 = -1.00 D of Corneal Astigmatism.

Example Recording:

Keratometry:

- OD: 42.00 @ 180 / 44.00 @ 90 (Mires clear) -> -2.00 D Astigmatism.

- OS: 42.00 @ 180 / 42.50 @ 90 (Mires distorted) -> Irregular.